## Slide 1: JOINT COMMISSION UPDATE 2014 Susan B. McLaughlin
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 2: COMING ATTRACTIONS
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 3: MEDICAL EQUIPMENT
Effective 7/1/14
Based on CMS changes to allow for varying maintenance strategies
Per manufacturers’ recommendations:
  Lasers
  Imaging & radiological equipment
  New equipment with no history
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 4: UTILITIES
Effective 7/1/14
Use manufacturer’s recommendations when insufficient history to support alternate strategies
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 5: DIAGNOSTIC IMAGING
Effective 7/1/14
13 EPs across 5 EC standards
  Safety
  Hazmat
  Medical Equipment
  Medical Equipment Maintenance
  Construction & Renovation
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 6: EC.02.01.01 SAFETY/SECURITY
EP 14: Hospitals providing MRI manage patient safety risks:
  Claustrophobia, anxiety, emotional distress
  Urgent or emergent medical care
  Implants, devices, or embedded foreign objects
  Ferromagnetic objects entering MRI environment
  Acoustic noise
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 7: EC.02.01.01 SAFETY/SECURITY
EP 16: Manages safety risks by:
  Restricting access to scanner room and area immediately preceding entrance
  These areas under supervision of trained MRI staff
  Posting signage at entrance to scanner:
    Potentially dangerous magnetic field
    Magnet always on (unless otherwise designed)
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 8: EC.02.02.01 HAZMAT
EP 17: Hospitals providing CT, PET, nuclear medicine
  Review of staff dosimetry at least quarterly by radiation safety officer or diagnostic medical physicist
  Staff exposure ALARA (as low as reasonably achievable) or below limits
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 9: EC.02.04.01 MEDICAL EQUIPMENT
EP 7: Identify quality control and maintenance activities to maintain the quality of images
  Frequency of activities
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 10: EC.02.04.03 MEDICAL EQUIPMENT MAINTENANCE
EP 15: Maintain quality of images
EP 17: Hospitals performing CT, annually, a diagnostic medical physicist:
  Measures radiation dose for specific CT scans
  Verifies dose displayed on CT is within 20% of displayed dose
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 11: EC.02.04.03 MEDICAL EQUIPMENT MAINTENANCE
EP 19, 20, 21, 22: Hospitals performing CT/MRI/Nuclear Medicine/PET, annually, a diagnostic medical physicist conducts performance evaluation of CT/MRI/Nuclear Medicine/PET equipment which includes specific imaging metrics.
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 12: EC.02.04.03 MEDICAL EQUIPMENT MAINTENANCE
EP 23: Hospitals performing CT/MRI/Nuclear Medicine/PET, annual performance evaluation by diagnostic medical physicist includes specific parameters
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 13: EC.02.06.05 CONSTRUCTION, RENOVATION, DEMOLITION
EP 4: Hospitals performing CT/Nuclear Medicine/PET, prior to installation of new or replacement equipment or modification of rooms, medical physicist conducts a “structural shielding design” to specify required radiation shielding.
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 14: EC.02.06.05 CONSTRUCTION, RENOVATION, DEMOLITION
EP 4: Hospitals performing CT/Nuclear Medicine/PET, after installation of imaging equipment or construction of rooms where radiation will be emitted or stored, a medical physicist conducts a radiation protection survey to verify adequacy of shielding.
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 15: CMS PROPOSAL EM
Add section 482.15: COP: Emergency Preparedness
Many proposals
  Some redundant to current standards
  Some significant new requirements
Comments closed 3/31/14
  ASHE submitted comments developed by task force
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 16: CMS PROPOSAL: NFPA 101 & 99 (2012)
http://www.ofr.gov/OFRUpload/OFRData/2014-08602_PI.pdf
Delete NFPA 99 Chapter 12: Emergency Management
Remove exemptions for 3 or fewer patients
Require smoke removal in anesthetizing locations
Comments open till 6/16/14
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 17: CMS/JC WAIVERS
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 18: CATEGORICAL WAIVERS
Application to CMS is not required
Written documentation stating that the waiver is being used
  “Additional Comments” on BBI
Present documentation to survey team
Meet all conditions of waiver
  See information from CMS and ASHE
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 19: CATEGORICAL WAIVERS
Medical Gas Master Alarms:
  Centralized computer system may substitute for one of the two required Category 1 alarm panels
Openings in Exit Enclosures:
  Existing openings to mechanical spaces permitted in exit enclosures if they are protected by fire-rated door assemblies
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 20: CATEGORICAL WAIVERS
Emergency Generators:
  Load bank time reduced to 1.5 hours
Door Locking:
  Special arrangements allowed for clinical and/or security needs & more than one delayed egress in exit path
Clean Waste & Recycling Containers
  96 gallons outside of hazardous area
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 21: CATEGORICAL WAIVERS
Suites:
  Permits exiting through another suite
  One of the two exits may be to an exit stair, exit passageway, or door to exterior
  Sleeping suites up to 10,000 square feet
Sprinkler Systems:
  Certain water flow devices test semiannually
  Fire pump no-flow test to monthly
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 22: CATEGORICAL WAIVERS
OR Humidity
  20-60%
Corridor Equipment without Time Restrictions
  Patient lifts
  Fixed furniture
Kitchen Equipment Open to Corridor
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 23: CATEGORICAL WAIVERS
Fireplaces
  Direct vent gas fireplaces in patient sleeping compartments
  Solid-fuel burning fireplaces in non-patient sleeping areas
Combustible Decorations
  Increases amount of wall space permitted to be covered
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 24: JOINT COMMISSION TOP TEN
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 25: 2013 JOINT COMMISSION TOP 10
LS.02.01.20 52%: Corridor Clutter
RC.01.01.01 52%
LS.02.01.10 48%: Penetrations; Fire Doors
EC.02.05.01 47%: Airborne Contaminants
IC.02.02.01 46%: Infections/Med Equipt.
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 26: 2013 JOINT COMMISSION TOP 10
LS.02.01.30 45%: Hazardous Areas; Suites
EC.02.03.05 45%: Fire Protection Testing
EC.02.06.01 39%: Gas Cylinders; Humidity
LS.02.01.35 36%: Sprinklers
MM.03.01.01 35%
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 27: TOP 10 #1
LS.02.01.20 EP 13: Corridor Clutter 52%
  Educate staff to risk
    Patient movement
    Staff movement
    Additional staff responding
    to emergency
    Joplin story
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 28: TOP 10 #4
EC.02.05.01 EP 6: Airborne Contaminants; Ventilation 47%
  Stained ceiling tiles
  Pressure differentials
  Air changes
  Tissue test is pre-screening
  Testing and balancing report (optional)
    When done (seasonal issues)
    Any specific requirements, such as doors closed
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 29: PRESSURE DIFFERENTIALS
ORs
Sterile Core
C-Section Rooms
Sterile Storage
Endoscopy****
Scope Cleaning
Bronchoscopy
Cath Lab
Interventional Radiology
Sterile Processing
Isolation Rooms
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 30: RECOMMENDATION (NOT regulation)
Measure pressure differentials AT LEAST annually
Verify AT LEAST quarterly
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 31: TOP 10 #4 continued
EC.02.01.05 EP 6 = CONDITION LEVEL DEFICIENCY
  If organization can repair process, LSC surveyor may review
  LCS surveyor may contact Central Office to discuss possibility of reducing CLD to SLD, with no change to finding
  Resolution must include area affected
    Balancing zone
    Ongoing assessment
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 32
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 33: HIGH LEVEL DISINFECTION SELF-CONTAINED UNITS (GUS, ETC.)
Focus on process and PPE
NO VENTILATION REQUIREMENTS
  General room ventilation
  Local exhaust hood
  Ductless fume hoods
  Replace filters per manufacturer
Storage to prevent contamination or damage
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 34: TOP 10 #5
IC.02.02.01: Reducing risk of infections associated with medical equipment, devices, and supplies. 46%
EP 2: When performing intermediate and high-level disinfection and sterilization of medical equipment.
  Scopes, contamination
  Immediate Threat to Life
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 35
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 36: TOP 10 #5 continued
Issues:
  Training and competency
  Failure to update/follow policy and procedure
  Inadequate supervision
CoP Infection Control 482.42
CoP Surgical Services 482.51(b)
Clinical Engineering scope inventory
  Oversight of cleaning per manufacturer recommendations
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 37: TOP 10 #8
EC.02.06.01 EP 1 & 13 39%
EP 1: Safe, suitable environment
  Unsecured oxygen cylinders
  Note: EP 5 – outdoor safety – compactor issues
EP 13: Ventilation, temperature, humidity
  Doors held open by air pressure
  Temperature: hot, cold calls
  Humidity >60% in OR. (Occasional spikes OK in rainy season per JC) (NOT CMS)
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 38
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 39: QUESTIONS?
smclaughlin@mslhealthare.com
847-420-3229
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE

## Slide 40: THANK YOU!
© 2014 MSL Healthcare Consulting
DO NOT DUPLICATE